## Slide 1: Recent Advances in the Textural Characterization of Porous Materials and Powders – Selected Aspects
Prof. Dr. Matthias Thommes
Institute of Separation Science and Technology
Department of Chemical and Biological Engineering (CBI)
University Erlangen – Nürnberg, Germany

## Slide 2
Founded in 1743, the Friedrich-Alexander University Erlangen-Nuremberg is today one of Germany's top comprehensive research universities with approximately 40,000 students.
.

## Slide 3
Physical Adsorption (Micro/Mesopore Analysis)
Liquid Intrusion/Mercury Porosimetry (Meso/Macropore Analysis)
Thermoporometry
Scattering Techniques (x-ray and neutron scattering techniques such as SAXS, SANS, XRD)
Spectroscopy (e.g., NMR based techniques)
Microscopy, (e.g.SEM, TEM), Digital Image Analysis etc..
Selected Experimental Techniques
Nanoporous Materials Characterisation
Combining gas adsorption with mercury porosimetry allows to assess a pore size range from ca. 0.3 nm up to > 400 m!

## Slide 4
Macropore Analysis:
Analogy between mercury intrusion/extrusion
and gas adsorption (pore condensation/evaporation)
Mechanism of mercury intrusion/extrusion can be considered thermodynamically equivalent to pore evaporation/condensation in gas adsorption → potential for advancing macropore size/pore network analysis by Hg Poroismetry, i.e. transfer of methodologies such as hysteresis scanning to mercury porosimetry
R. Guillet-Nicolas, K.A. Cychosz, F. Kleitz, M. Thommes, New. J. Chem. 40 (2016) 4251
F. Porcheron, P. Monson, M. Thommes, Langmuir 20 (2004) 64
F. Porcheron, M. Thommes, R. Ahman, P. Monson, Langmuir 23 (2007) 3372.
N2/77K
Ads/Des iostherms of N2 and Hg in
ordered mesoporous silica KIT-6
KIT-6 silica
Wetting
Hg/293 K)
Non-wetting

## Slide 5
Adsorption isotherms of Argon (87 K) and Water (298 K) into
fully hydrophobized mesoporous silica (pore size: 9.1 nm)
H20 (296K)
H20
Ar (87K)
●Water adsorption isotherms obtained from thermodynamic analysis
of water intrusion
J. Söllner and M. Thommes, to by published,
Jakob Söllner

## Slide 6
Textural Characterization Techniques
C. Schlumberger, M. Thommes, Advanced Materials Interfaces (2021), 5. 2002181
P. Klobes, K. Meyer, G. Munro, NIST, Gaithersburg, MD 2006, p.960

## Slide 7
Characterization of Porous Materials and Powders
● Physical Adsorption – Micro- Mesopore Analysis
- Major progress during the last two decades in understanding
the adsorption and phase behaviour of fluids in pores
● Liquid Intrusion (e.g. Mercury Porosimetry) – Macropore Analysis
- Recent advances in the understanding of phase behavior of
nonwetting fluids under confinement
Some Challenges:
● Assessment of Surface Chemistry: i.e. only possible by combining
complimentary techniques such as adsorption, liquid intrusion,
and spectroscopy
● Textural characterization of materials immersed in liquid phase
- Development of novel methdologies based on:
- NMR Relaxometry
- High frequency conductivity measurements
- Electroacoustics

## Slide 8: Physical Adsorption Characterization of Nanoporous Materials
Surface Area
Pore Size/Volume Distribution/Porosity
Pore Geometry, Structure of Pore Network
Surface Properties

## Slide 9: New IUPAC Recommendations (2015) Updated Classification of Physisorption Isotherms
Nanopores: pore widths ≤ 100 nm
Within the range of nanopores:
Micropores < 2 nm
Ultramicropore < 0.7 nm
Supermicropore 0.7 – 2nm
Mesopore: 2- 50 nm
Macropores: > 50nm
,
Classification represents subcritical adsorption isotherms on rigid adsorbents
ultramicropore
supermicropore
non-porous; het.surf
non-porous, het surf.
mesoporous
mesoporous
microporous
or–mesoporous
nonporous-homog.
surface
M. Thommes, K. Kaneko, A.V. Neimark, J.P. Olivier, F. Rodriguez-Reinoso J. Rouquerol, K.S.W. Sing, Pure Appl. Chem., 2015, 87, 1051.
New
New

## Slide 10: Mesopores (Type IV isotherm): MultilayerAdsorption, Phase Behavior, and Hysteresis
adsorption branch: delayed condensation due to metastable pore fluid
desorption branch: equilibrium
M. Thommes , In: Nanoporous Materials, Science & Engineering
(Lu G.Q. & Zhao X.S., eds), Chapter 11, Imperial College Press, 2004

## Slide 11
Adsorption and desorption of a fluid in a mesopore
at temperature T1 and T2 (T1 < T2)
C. Schlumberger, M. Thommes, “Characterization of hierarchically ordered porous materials
by physisorption and mercury porosimetry- (invited) A Tutorial Review”,
Advanced Materials Interfaces (2021), 6, 2002181
Vapour-Liquid Phase Transition can be understood as homogeneous
nucleation process

## Slide 12
Effect of confinement on pore condensation
and hysteresis
Schematic phase diagram of pore and bulk fluid applicable to fluids such as Ar and N2
confined in mesoporous oxides and carbons,
M. Thommes, Physical adsorption characterization of ordered and amorphous mesoporous
materials. In Nanoporous Materials: Science and Engineering, vol. Volume 4 of Series on
Chemical Engineering (2004) 317–364;
M. Thommes, C. Schlumberger, Annual Review of Chemical and Biomolecular Engineering (2021) 12, 137

## Slide 13: CO2 adsorption and phase behavior in ordered micro- mesoporous CMK 3 carbon – Experiments over a wide range of temperatures and pressures
Vapor-solid
QSDFT Pore Size Analysis
from N2 (77K) and Ar (87K) isotherms
5 nm
185 K
195
220
240 K
273K
● No hysteresis is observed over a wide range of temperatures
→ shift of the hysteresis/pore critical temoerature in the order of  65K!!!!
1nm
● Capillary condensation with hysteresis occurs below CO2 bulk triple point but not at 195 K (dry ice/aceton temperature)
→ shift of pore triple point region to lower temperaures, i.e. the downward shift is > 6 K but < 20 K
273K
240 K
220K
210 K
Ttr, CO2 : 216.58 K
Tc, CO2 : 304.18 K
Dantas, Cychosz, Thommes*, Neimark* Langmuir, 2019, 35, 11291

## Slide 14
Surface Area Analysis:
BET Area Determination

## Slide 15: New IUPAC Recommendations (2015) Updated Classification of Physisorption Isotherms
Nanopores: pore widths ≤ 100 nm
Within the range of nanopores:
Micropores < 2 nm
Ultramicropore < 0.7 nm
Supermicropore 0.7 – 2nm
Mesopore: 2- 50 nm
Macropores: > 50nm
,
Classification represents subcritical adsorption isotherms on rigid adsorbents
ultramicropore
supermicropore
non-porous; het.surf
non-porous, het surf.
mesoporous
mesoporous
microporous
or–mesoporous
nonporous-homog.
surface
M. Thommes, K. Kaneko, A.V. Neimark, J.P. Olivier, F. Rodriguez-Reinoso J. Rouquerol, K.S.W. Sing,
Pure Appl. Chem., 2015, 87, 1051.

## Slide 16: Determination of the specific surface area of solids by gas adsorption – BET (Brunauer, Emmett and Teller method)
●Analysis of adsorption isotherm in a defined range of pressures (i.e. “classical “ relative pressures range : P/P0:
0.05 – 0.3)) with the BET equation allows one to determine the (statistical ) monolayer of adsorbate
coupled with the known cross-sectional area of an adsorbate atom/molecule the surface area can be determined
BET Equation
●The Brunauer-Emmett-Teller (BET) method continues to be the most widely used procedure for evaluating the surface area of porous and finely-divided materials from physical adsorption isotherms
Published in 1938

## Slide 17: Reference method: Gas adsorption
Reliable Surface Area Assessment by Adsorption (BET method)
  Overestimation of surface area possible with nitrogen (approx. 20-25 %) [1,2]
Solution: Argon (87 K) adsorption [1-3]
[1] C. Schlumberger, …G Reichenauer, M. Thommes et al., Reliable surface area determination of powders and meso/macroporous materials: Small-angle X-ray scattering and gas physisorption, Microporous and Mesoporous Materials. (2022) 111554.
[2] M. Thommes, C. Schlumberger, Characterization of Nanoporous Materials, Annual Review of Chemical and Biomolecular Engineering. 12:137-162 (2021) 137–162 and references therein
[3] Thommes M et al. 2015. Physisorption of gases, with special reference to the evaluation of surface area and pore size distribution (IUPAC Technical Report). Pure and Applied Chemistry. 87(9–10):
Surface area:
Standard adsorptive for many years: Nitrogen (77 K)
  Problem: Specific interaction with polar surface sites due to quadrupole moment
  Cross-sectional area is depending on surface chemistry
References

## Slide 18: Surface Area Analysis
Benchmark and surface area data on silica model materials
without micro-and narrow mesopores
[1] Schlumberger C , Reichenauer, M Thommes et al. 2021. Reliable Surface Area Determination of of powders and meso/macroporous materials : Small Angle X-ray Scattering and Gas Physisorption, Microporous and Mesoporous Materials. (2022) 111554
[2] Thommes, M. et al. 2015. Physisorption of gases, with special reference to the evaluation of surface area and pore size distribution (IUPAC Technical Report), Pure and Applied Chemistry. 87
 | S (Ar, 87 K) / m2 g-1 | S (N2, 77 K) / m2 g-1
CPG ERM-FD-121 | 98.6 | 118.6
Silica, 12 nm | 162.6 | 202.5
Silica, 200 nm | 15.6 | 18.9
Silica, 80 nm | 47.1 | 57.5
Silica, 120 nm | 23.6 | -
Silica, 1 µm | 2.6 | -
  Systematic SAXS measurements independenly confirm accurate surface area determination with
  Ar (87 K), while N2 (77K) is unreliable for assessing
  of surface with polar surface functionality
  (errors up to 25- 30 %)1,2
Argon 87 K
on nonporous Silica
SAXS
Notes: Cooperation with Anton Paar and ZAE Würzburg, validation of surface area with SAXS

## Slide 19: Problem: Application of the BET Method for Determination of the BET Area of Microporous Materials
Problem: BET equation is in a strict sense not applicable to microporous materials  BET Area (apparent surface area)
Application of BET equation in “classical range”, i.e. relative pressure range 0.05 to 0.3:
  Negative y-intercept
  Negative C constant
  Less than optimal correlation coefficient
BET area: 1008 m2/g,

## Slide 20: Determination of the BET Area of Microporous Materials
Find the linear range of the BET plot for microporous materials in a way that reduces any subjectivity
J. Rouquerol et al. (Stud. Surf. Sci. Catal. 2007, 160, 49-56) suggested the following
main criteria:
  (1) The quantity of C must be positive. Any negative intercept on the ordinate of the BET plot is an indication that one is outside the valid range of the BET equation.
  (2) The application of the BET equation should be limited to the pressure range where the term n(P0-P) or n(1-P/P0) continuously increases with P/P0. All data points above the maximum of the plot should be eliminated.
  (3)The pressure finally corresponding to the monolayer capacity, nm, should enter the pressure range for the calculation.

## Slide 21: Procedure to Find the Linear BET Range for Microporous Materials
  Improves reproducibility of BET area determination of microporous materials
  Improves the comparison of BET area data between labs-
microporous material
BET area = 1167 m2/g
maximum at P/P0 = 0.07, apply BET equation to data points below that
positive C constant, good correlation coefficient

## Slide 22: IUPAC Report 2015 Surface Area Determination
● ●The Brunauer-Emmett-Teller (BET) method continues to be the most widely used procedure for evaluating the surface area of porous and finely-divided materials
●Under certain carefully controlled conditions, the BET-area of a non-porous, macroporous or a mesoporous solid (i.e. giving a well-defined Type II or a Type IVa isotherm) can be regarded as the probe accessible surface area
● In the presence of micropores (ie. Type I isotherms and combinations of Types I and II or Types I and IV isotherms) the application of the BET method
only leads to an apparent surface area (i.e BET Area) which serves as a useful “fingerprint” of the adsorbent.
In this case the linear range of the BET plot may be difficult to locate. A useful procedure suggested by Rouquerol et. al permits to overcome this difficulty and improves the reproducibility of the method.

## Slide 23: Development of methodology for a reliable textural characterization of materials immersed in the liquid phase by NMR relaxamotry
 Relaxation of liquid on the particle surface is
much faster than in the bulk liquid
Spin-spin relaxation (T2):
C. Schlumberger
Particle
Notes: Surface area assessment is based on the fact that liquid near the particle surface shows a shorter relaxation time compared to the bulk liquid.

## Slide 24: Surface Area Assessment by NMR Relaxometry
NMR relaxation of nonporous silica materials in water
 | Silica 1 µm | Silica 480 | Silica 200 | Silica 140 | Silica 80 | Silica 12
S (BET, Ar) / m2 g-1 | 2.6 | 6.9 | 13.7 | 24.2 | 40.7 | 162.6
S (NMR, T2) / m2 g-1 | 2.2 | 7.1 | 13.6 | 29.2 | 47.3 | 161.4
Surface area of silica model
materials (nonporous silica / water)
          Good agreement between NMR T2 surface area and BET surface area
M. Thommes, C. Schlumberger, Characterization of Nanoporous Materials, Annual Review of Chemical and Biomolecular Engineering. 12:137-162 (2021) 137–162
Schlumberger, C., Sandner, L., Michalowski, A., Thommes, M. Reliable surface area assessment of wet and dry nonporous and nanoporous particles: Nuclear magnetic resonance relaxometry and gas physisorption, under review
Good correlation between T2 relaxation time and surface area
Increase in surface area

## Slide 25: Surface Area Assessment by NMR Relaxometry
Good agreement between NMR T2 surface area and BET surface area for carbon blacks
 | Carbon Black N774 | Carbon Black N375
S (BET, Ar) / m2 g-1 | 26.1 | 82.2
S (NMR, T2) / m2 g-1 | 27.7 | 69.9
Surface area of carbon blacks (Carbon Black N134/THF)
[1] D. Fairhurst et al., Relaxation NMR as a tool to study the dispersion and formulation behavior of nanostructured carbon materials, Magnetic Resonance in Chemistry. 54 (2016) 521–526.
[2] Schlumberger, C., Sandner, L., Michalowski, A., Thommes, M. Reliable surface area assessment of wet and dry nonporous and nanoporous particles: Nuclear magnetic resonance relaxometry and gas physisorption, under review
Spin-spin relaxation of carbon blacks
in tetrahydrofuran (THF) [1]
Good correlation between T2 relaxation time and surface area
Increase in surface area

## Slide 26
Pore Size Analysis

## Slide 27: New IUPAC Recommendations (2015) Updated Classification of Physisorption Isotherms
Nanopores: pore widths ≤ 100 nm
Within the range of nanopores:
Micropores < 2 nm
Ultramicropore < 0.7 nm
Supermicropore 0.7 – 2nm
Mesopore: 2- 50 nm
Macropores: > 50nm
,
Classification represents subcritical adsorption isotherms on rigid adsorbents
ultramicropore
supermicropore
non-porous; het.surf
non-porous, het surf.
mesoporous
mesoporous
microporous
or–mesoporous
nonporous-homog.
surface
M. Thommes, K. Kaneko, A.V. Neimark, J.P. Olivier, F. Rodriguez-Reinoso J. Rouquerol, K.S.W. Sing, Pure Appl. Chem., 2015, 87, 1051.
New
New

## Slide 28
● Pore Size/Volume Characterization of microporous materials with polar surface functionality, e.g. zeolites, MOFs
 Argon 87 K adsorption
● Characterization of ultramicroporous carbons
   CO2 adsorption at 273K
● Low surface area materials
 Krypton adsorption at 77 K for BET analysis
IUPAC (2015): Key Recommendations
for Characterization of Porous Materials
Choice of Adsorptive:

## Slide 29: High Resolution Argon (87 K) Adsorption
mordenite
soc-MOF
Thommes, M. Textural Characterization of Zeolites and Ordered Mesoporous Materials by Physical Adsorption in: Introduction to Zeolite Science and Practice, 2007, p. 495-523
Thommes, M.; Cychosz, K.A. Adsorption 2014, 20, 233-250
soc-MOF
Mordenite
ZSM-5
→ Argon 87 K adsorption allows to resolve pore width differences on a 0.1 nm level

## Slide 30: Argon (87 K) Adsorption in Zeolites and MOFs
Mordenite
→ Argon 87 K adsorption allows to resolve pore width
differences on a 0.1 nm level
NLDFT Pore Size Analysis
ZSM-5
Argon 87 K Adsorption

## Slide 31: Pore Size/Volume Analysis of Nanoporous Materials
Classical, macroscopic methods
  Micropores (< 2 nm): e.g. Dubinin-Radushkevitch, Horvath-Kawazoe (HK), Saito-Foley (SF), comparison plot methods (t-method, alpha-s method)
  Note: Application of conventional t-method and other comparison plot methods is not straightforward for Micro-mesoporous materials with narrow mesopores
  Meso- and Macropores (> 2 nm): e.g. Kelvin equation based methods such as Barrett, Joyner, Halenda (BJH) or Broeckhoff-de Boer (BDB), KJS (Kruk,Jaroniec, Sayari), VBS (Villarroel-Rocha, Barrera, Sapag)
  Note: Classical methods (if not corrected) understimate pore sizes of pore width < 10 nm up to 25-30%
Modern, microscopic methods based on statistical mechanics describe configuration of adsorbed molecules on a molecular level (IUPAC 2015)
  Both Micro- and Mesopore size range: e.g. Density Functional Theory (DFT), Molecular Simulation
An accurate pore size analysis over the complete pore size range can be
 performed by a single method.

## Slide 32
Pore Size Analysis of SBA-15 Silica
Kelvin/BJH
NLDFT
Equilib. transition
delayed
condens..
M.Thommes, R. Guillet-Nicolas, K.A Cychosz, Physical Adsorption Characterization of Mesoporous Zeolites. In: Mesoporous Zeolites: Preparation, Characterization and Applications, Ed. J. Garcia-Martinez, K. (Eds) 2015, Wiley, p.
Landers, G. Gor, A.V. Neimark, Application of Density Functional Theory Methods for Characterization of Porous Materials., Colloids Surfaces A, 2013

## Slide 33: Characterization of Hierarchically Structured ZSM-5 Zeolite by Argon (87 K) Adsorption and NLDFT Analysis
Zhang, X.; Liu, D.; Xu, D.; Asahina, S.; Cychosz, K.A.; Agrawal, K.V.; Al Wahedi, Y.; Bhan, A.; Al Hashimi, S.; Terasaki, O.; Thommes, M.; Tsapatsis, M,. Science 2012, 336, 1684-1687
Unique house-of-cards arrangement of nanosheets provides a hierarchical pore net
work.

## Slide 34: Characterization of Hierarchically Structured Nanoporous Carbon for Supercapacitor Applications by combining N2, Ar and (77, 87K) and CO2 (273K) Adsorption
NLDFT Analysis : slit/cylinder mod.
X.Zhu., S. Murali, M. D. Stoller, K. J. Ganesh, W. Cai, P.J. Ferreira, A.Pirkle, R.M. Wallace, K. A. Cychosz, M.Thommes, D.Su, E.A. Stach and R. S. Ruoff.
Science 2011, 332, 1537-1541
,

## Slide 35
Assessing Pore Network Characteristics
by Advanced Adsorption Methodologies
Key:
● Understanding of Adsorption and Phase Behavior
of Fluids in Pores (confined geometry effects)
● Understanding contributions to Hysteresis

## Slide 36: IUPAC (2015) Classification of Hysteresis Loops
M. Thommes, K. Kaneko, A.V. Neimark, J.P. Olivier, F. Rodriguez-Reinoso
J. Rouquerol, K.S.W. Sing, Pure Appl. Chem., 2015, 87, 1051.
Ultramicropores
Heterogeneous Planar Surfaces
Heterogeneous Planar Surfaces
Narrow Mesopores
“Ideal” Planar Surfaces
Mesopores and Micropores
Supermicropores
Large Mesopores
Cylindrical Pores
Pore Networks and Ink Bottle Pores
Disordered pore structures
Micro- and mesoporous adsorbents
Open and restricted mesopores

## Slide 37
Correlation between
pore structure and
underlying adsorption
mechanism
Adsorption Hysteresis
Pore Network Analysis by Physisorption:
Requires understanding
of hysteresis
Thommes, M.; Cychosz, K.A. Adsorption 2014, 20, 233-2
Thommes, M. Schlumberger C,Annu. Rev. Chem. Biomol. Eng. 2021. 12:137–62

## Slide 38
Desorption/Evaporation Mechanisms:
Pore Blocking vs Cavitation
C. Schlumberger, M. Thommes, “Characterization of hierarchically ordered porous materials
by physisorption and mercury porosimetry- (invited) A Tutorial Review”, Advanced Materials Interface (2021), 6, 2002181

## Slide 39: Combined Micro- and Mesopore Size Analysis of Hierarchically Structured KLE Silica
N2 (77 K) adsorption isotherm
Mesopore size:
  N2 adsorption = 13.9 nm
  TEM = ~13 nm
  SAXS = 13.8 nm
NLDFT pore size distribution (spherical mesopores, cylindrical micropores from adsorption branch
Thommes, M.; Smarsly, B.; Groenewolt, M.; Ravikovitch, P.I.; Neimark, A.V. Langmuir 2006, 22, 756
Cavitation 
1.3 nm
13.9 nm
Excellent agreement between SAXS, TEM, and NLDFT approach
KLE
Pore diameter (Angström)
N2/ 77 K

## Slide 40
Pitfall: Effect of Cavitation on Nitrogen Sorption in Disordered Alumina Catalyst Support
Artifact
Desorption PSD
Cavitation→
→ The typical “step-down” of the desorption branch at ca. 0.4 for nitrogen adsorption is indicative of cavitation , i.e. which occurs because large mesopores are only accessible through narrow mesopores (pore withd < ca. 6 nm) or micropores.

## Slide 41: “Physisorption Toolkit “: Pore Network Analysis
(1) Application of advanced DFT methods for analysis of hysteresis loop
- checking consistency of PSD’s from Adsorption and Desorption Branches
→ can detect whether pore/pore network is open(unrestricted) or if restricted
pores exist which could give rise to either pore blocking or cavitation
controlled evaporation(desorption)
(2) Adsorption measurements with different adsorptives (e.g. argon, nitrogen)
→ Differentiating between pore blocking and cavitation in restricted pore
systems
(3) (Scanning of hysteresis loop
→ Differentiating between pore blocking and cavitation in restricted pore
systems
(4) Adsorption measurements with a given adsorptive at different temperatures
→ changing thermodynamics of confined pore phase
For details: C. Schlumberger, M. Thommes, “Characterization of hierarchically ordered porous materials
by physisorption and mercury porosimetry- (invited) A Tutorial Review”,
Advanced Materials Interface (2021), 6, 2002181

## Slide 42
Advanced Pore Network Analysis Approach
C. Schlumberger, M. Thommea”, Advanced Materials Interface (2021), 6, 2002181

## Slide 43: Advanced Pore Network Analysis by “Scanning” of the Hysteresis Loop
Cimino, R.; Cychosz, K.A.; Thommes, M.; Neimark, A.V. Colloids Surfaces A 2013, 437, 76-89
Type H1 Hysteresis: Pores evaporate independently from one another (e.g. SBA-15 silica)
Type H2 Hysteresis: Evaporation of certain pores depends on the state of neighboring pores  Network effects/percolation
Pioneered by Everett - 1967, De Boer - 1958
Pore Network - Type H2 hysteresis
Quantitative analysis of scanning curves allows one to obtain information about pore connectivity

## Slide 44
Hysteresis Loop Scans and Pore Network Structure
From: S. A. Kube, K. Turke, R. Ellinghaus, D. Wallacher, M. Thommes,B. M. Smarsly, Langmuir 2020, 36, 11996.
R. Cimino; K.A. Cychosz, M.Thommes, A.V. Neimark, A.V. Colloids Surfaces A 2013, 437, 76-89

## Slide 45
KIT-6 silica
Scanning and Pore Network Analysis
KLE-silica
Cimino, R.; Cychosz, K.A.; Thommes, M.; Neimark, A.V. Colloids Surfaces A (2013)(, 437, 76-89
J. Söllner, M. Thommes, manuscript in preparation, (2023)
Rasmussen, C.J.; Vishnyakov, A.; Thommes, M.; Smarsly, B.M.; Kleitz, F.; Neimark, A.V. Langmuir (2010)
Cavitation
-Pore Neck
-Pore Cavities
- Pore connecticity
Information:

## Slide 46
K.A. Cychosz, K.A., R.Guillet-Nicolas, J. Garcia-Martinez and M. Thommes*, “Recent advances in the textural characterization of hierarchically structured nanoporous materials“ (invited) Chem. Soc. Rev. 46 ,389 (2017)
Assessing pore network structure/connectivity
Example: hierarchical micro-mesoporous zeolites

## Slide 47: Argon (87 K) Adsorption and NLDFT Pore Size Analysis of Mesoporous Y Zeolites
Garcia-Martinez, J.; Xiao, C.; Cychosz, K.A.; Li, K.; Wan, W.; Zou, X.; Thommes, M. ChemCatChem 2014, 6, 3110-3115
A)
C)
D)
B)

## Slide 48: Advanced Characterization Hierarchically Structured Mesoporous Y-Zeolite (C16)
Collaboration with University of Stockholm, X. Zhou
(1) Advanced image analysis: transmission electron microscopy (TEM), including high-resolution imaging and electron tomography (ET); three-dimensional rotation electron diffraction
  Directly revealed the connectivity of micro- and mesopores in the mesostructured zeolite
(2) Advanced gas adsorption: high resolution argon adsorption at various temperatures;hysteresis scanning; application of advanced NLDFT methods
  Quantitative determination of intracrystalline micro- and mesopores volumes and differentiation between truly accessible (open) mesopores and those that are constricted
Garcia-Martinez, J.; Xiao, C.; Cychosz, K.A.; Li, K.; Wan, W.; Zhou, X.; Thommes, M. ChemCatChem 2014, 6, 3110-3115
Highlight: Thomas, J.M.; Leary, R.K. Angew. Chemie. Int. Ed. 2014, 53, 2-4

## Slide 49: Application of “Physisorption Toolkit” : Changing of state of pore fluid to visualize pore network connectivity
Garcia-Martinez, J.; Xiao, C W.; Zou, X.; Thommes, M. ChemCatChem 2014, 6, 3110-3115
A)
C)
D)
B)
N2 (77 K) and Ar (87 K)
NLDFT PSD
Ar (77 K)
Ar (77 K)
Ar (65 K)
open -->
restricted -->
Detailed structural information from hysteresis scanning
4.2 nm
1.2 nm
 More than 65% of pores lead to micropores
Tr(Ar): 83.8 K

## Slide 50: Confined Geometry Effects: Adsorption and Phase Behavior of Fluids in Mesopores
pore W1 > pore W2
Ar/MCM-41
Ar 87K/
MCM-41
Morishige et al.,1997
Kruk et al, 2000
Schematic phase diagram of pore and bulk fluid,
applicable to Ar and N2confined in
mesoporous oxides and carbons ,
IV(a)
IV(b)
M. Thommes et, al, Applied, Surface Science, 2002)
pore W1 > pore W2

## Slide 51
Summary
● Physical adsorption and mercury porosimetry are powerful tool s for the
characterization of porous materials
● Advanced, high resolution physical adsorption experiments (e.g,
scanning of hysteresis etc) coupled with state-of-the-art data
reduction (DFT and molecular simulation based methods
 allows one to obain reliable surface and pore size/porosity information
and explore important aspects of pore connectivity
  What is needed for a comphensive and accurate textural characterization of porous materials and powders with complex and hierarchical pore structures:
  ● Development of new procesures/methodologies based on the combination of adsorption and mercury porosimetry with complimentary experimental techniques such as in-situ scattering (e.g., SAXS,SANS) electron tomography and reconstruction, spectroscopy (such as NMR-based methods)
  ● Further work needed for deriving of a unified framework for textural characterisation of functionalised, porous materials and particles in the wet and dry state.

## Slide 52: Acknowledgement
Institute of Separation Science and Technology (University Erlangen-Nuremberg)
Federal Ministry for
Economic Affairs and Energy
CRC 1411
CRC 1452

## Slide 53
Macropore Analysis

## Slide 54: Limitation of Adsorption Characterization Nitrogen Sorption at 77 K into Mesoporous TiO2
6 nm
10 nm
30 nm
100 nm
H. Kueppers, B. Hirthe, M.Thommes, G.I.T, 3 (2001) 110
Gas Adsorption: reliable pore size analysis only for pores with diameter < 50 nm
200nm

## Slide 55
Pore size/volume analysis over a wide range from 4 nm up to 400 m
 “Fast” and highly reproducible measurements
Texture (Pore structure,Tortuisity, Permeability)
Surface area
Particle size distribution
Skeletal and Bulk Density
Washburn Equation:
Ph r = - 2cos

non-wetting
Macropore Analysis: Mercury Porosimetry

## Slide 56: Macro-Mesopore Size Distributon of Hierarchical Macro-Mesopore Silica Monolith for Chromatographic Applications
M. Thommes , R. Skudas,K.K. Unger, et al.J. Chromatogr. A, 2008,1191, 57

## Slide 57: Pore Network Analysis of Pore Systems with Hierarchical Macro-Mesopore Porenetwork
Studies of Hg intrusion/extrusion into porous glasses (Vycor and CPG)
by combining experiments and advanced molecular simulation confirm:
  Hysteresis in gas adsorption and Hg porosimetry has the same origin
  Entrapment is of kinetic origin, i.e., due to fragmentation (snap-off) of liquid during extrusion
F. Porcheron , P. Monson,, M. Thommes
et al., Langmuir 2004,20, 64
F. Porcheron, R. Ahmad, M. Thommes,
P. Monson. Lanngmuir, 2007, 23, 3372
M. Thommes , R. Skudas,K.K. Unger, et al.
J. Chromatogr. A, 2008,1191, 57
Further confirmed by Hg porosimetry on silica monoliths
No entrapment
Extrusion → cavitation

## Slide 58: Mercury Porosimetry into MCM-41 and SBA-15 Silica
Some selected examples:
- A. Lind, C. du Fresne Von Hoenesche, J.H. Smått, .; M. Linden;
K.K. Unger, Microporous Mesoporous Mater. 2003, 66, 219.
- M. Hartmann, A. Vinu, A. Langmuir 2002, 18, 8010.
- A. Vinu, V. Murugesan, W. Bohlmann, M. Hartmann, . J. Phys.Chem. B 2004
108, 11496.
- A. Galarneau, B. Lefèvre,† H. Cambon,† B. Coasne,† S. Valange, Z. Gabelica, J.-P. Bellat,
and F. Di Renzo*, J. Phys. Chem. C 2008, 112, 12921–12927
 These studies highlighted the effects of mechanical stability of
mesoporous molecular sieves (mainly with MCM-41, SBA-15 silica )
combined with surface roughness effects on mercury intrusion/extrusion

## Slide 59: Mercury Intrusion/Extrusion in ordered mesoporous silica (KIT-6)
First reversible mercury intrusion/extrusion experiments into ordered mesoporous 3D silica ( KIT-6 ) NO fracture
For more details: R. Guillet-Nicolas, R. Ahmad, K.A. Cychosz, F. Kleitz, M. Thommes, New J. Chem. 2016,

## Slide 60: Equivalency of Gas Sorption (pore condensation /evaporation) and Hg Porosimetry (extrusion/intrusion)
(1) Symmetry in Experiments:
Lowell and Shields (1981), Seaton and Murray (1999).
Correlation between the hydraulic pressure Ph (required to force mercury into pores) to the relative pressure P/P0, exerted by the liquid Hg with radius of curvature r:
Ph = (RT/VM) ln(P/P0)
(2) Symmetry in Theory:
Lattice gas model (Porcheron, Monson, Thommes ) exhibits a symmetry that provides a direct relationship between I intrusion/extrusion of non wetting fluid with adsorption/desorption isotherms for a wetting fluid
F. Porcheron , P. Monson,, M. Thommes
et al., Langmuir 2004,20, 64
F. Porcheron, R. Ahmad, M. Thommes,
P. Monson. Lanngmuir, 2007, 23, 3372

## Slide 61
Macropore Analysis:
Analogy between mercury intrusion/extrusion
and gas adsorption (pore condensation/evaporation)
Mechanism of mercury intrusion/extrusion can be considered thermodynamically equivalent to pore evaporation/condensation in gas adsorption → potential for advancing macropore size/pore network analysis by Hg Poroismetry, i.e. transfer of methodologies such as hysteresis scanning to mercury porosimetry
R. Guillet-Nicolas, K.A. Cychosz, F. Kleitz, M. Thommes, New. J. Chem. 40 (2016) 4251
F. Porcheron, P. Monson, M. Thommes, Langmuir 20 (2004) 64
F. Porcheron, M. Thommes, R. Ahman, P. Monson, Langmuir 23 (2007) 3372.
N2/77K
Ads/Des iostherms of N2 and Hg in
ordered mesoporous silica KIT-6
KIT-6 silica
Wetting
Hg/293 K)
Non-wetting

## Slide 62
Summary
● Physical adsorption and mercury porosimetry are powerful tool s for the
characterization of porous materials
● Advanced, high resolution physical adsorption experiments (e.g,
scanning of hysteresis etc) coupled with state-of-the-art data
reduction (DFT and molecular simulation based methods
 allows one to obain reliable surface and pore size/porosity information
and explore important aspects of pore connectivity
  What is needed for a comphensive and accurate textural characterization of porous materials and powders with complex and hierarchical pore structures:
  ● Development of new procesures/methodologies based on the combination of adsorption and mercury porosimetry with complimentary experimental techniques such as in-situ scattering (e.g., SAXS,SANS) electron tomography and reconstruction, spectroscopy (such as NMR-based methods)
  ● Further work needed for deriving of a unified framework for textural characterisation of functionalised, porous materials and particles in the wet and dry state.

## Slide 63: Acknowledgement
Institute of Separation Science and Technology (University Erlangen-Nuremberg)
Federal Ministry for
Economic Affairs and Energy
CRC 1411
CRC 1452

## Slide 64: Institute of Separation Science and Technology (TVT)
Adsorption and Characterisation of Porous Materials
Advanced Separation Processes
Adsorption fundamentals
Gas/energy storage and separation applications
Novel approaches for characterisation of surface and pore structure
Breakthrough measurements
Preparative chromatography
Membrane processes
Gas adsorption
Modelling and optimisation-based process design
Hysteresis scanning
Hierarchical mesoporous zeolite Y
MW2
Resolution
MW1
300
3000
300
3000
10
2
6
High-resolution chromatography
Hybrid SSR/racemisation process
NMR relaxation
Structure-
Property-Performance-Relationship
Process
Porous materials
Prof. Thommes
(Head of Institute )
Prof. Kaspereit
Chromatography
R
S
R↔S
R, S